# Mabtech IRIS™ 2/ASTOR™ 2 system specifications

# Introduction

Mabtech IRIS™ 2 FluoroSpot/ELISpot/FociSpot reader and Mabtech ASTOR™ 2 ELISpot reader use the software Mabtech Apex™ for exact spot center determination using RAWspot™ technology.

Both readers were developed by Mabtech AB in Nacka Strand, Sweden, whose quality management system complies with ISO 9001:2015 (International Organization for Standardization) and ISO 13485:2016. The readers are CE-marked in accordance with the EU Directive on Electromagnetic Compatibility (EMC, 2014/30/EU) and the Low Voltage Directive (LVD, 2014/35/EU). They are also compliant with RoHS (Restriction of Hazardous Substances), REACH (Registration, Evaluation, Authorisation and Restriction of Chemicals), WEEE (Waste Electrical and Electronic Equipment), FCC (Federal Communications Commission), and ICES (Interference-Causing Equipment Standard). The readers are intended for research use only (RUO) and not for use in diagnostic procedures. However, the technical performance of Mabtech IRIS™ 2 and Mabtech ASTOR™ 2 enables their use in regulated laboratories following user qualification.

# Mabtech IRIS™ 2 and Mabtech ASTOR™ 2 validation

Mabtech IRIS™ 2 and Mabtech ASTOR™ 2 are assembled, validated, and calibrated at Mabtech’s headquarters in Nacka Strand, Sweden, and distributed to end users worldwide. All readers are calibrated to match the light intensity of reference readers and given unique machine default exposure for each LED. These individually set machine default exposure values are then used by the software Mabtech Apex™  when reading ELISpot, FluoroSpot, and FociSpot plates. A validation document describing the validation results is supplied with each reader.

Installation Qualification and Operational Qualification (IQOQ)

Qualified Mabtech technicians install Mabtech readers. Installation includes reader set-up, verification of XY table alignment, and operational qualification of optics and light sources. The operational qualification is made using a plate produced at Mabtech containing spots with all fluorophores used in our kits, one for each filter in Mabtech IRIS™ 2: LED380, LED490, LED550, and LED640. In addition, standard ELISpot plates are utilized for qualification of the White LED used in ELISpot and FociSpot analysis in Mabtech IRIS™ 2 as well as Mabtech ASTOR™ 2.

Upon installation, the Mabtech technician makes sure that the plate generates comparable spot counts for White LED and LED380-, LED490-, LED550- and LED640-excited spots as it did when the reader was assembled and approved at Mabtech headquarters (<5% CV between plate readouts). This ensures that the reader has not suffered from any damage during transport and that the instrument matches reference readers at Mabtech headquarters. All procedures are documented in an installation qualification and operational qualification document. In addition to installation and operational qualification, Mabtech provides a general introduction to the reader and hands-on training with customer staff.

Performance Qualification

Mabtech offers Performance Qualification (PQ) plates used to ensure that Mabtech ASTOR™ 2 and Mabtech IRIS™ 2 perform according to specifications. With the PQ plate, Mabtech Apex™ software conducts a series of quantitative and qualitative tests that can identify any incorrect or inconsistent operation of the reader. The data obtained is compared to baseline levels set for camera focus, LED light intensities, and spot counting using the RAWspot™ algorithm. Each PQ plate comes with a unique serial number, which is registered in Mabtech Apex™, ensuring correct comparison. A report is automatically generated in non-modifiable PDF format, and the results are saved on the computer.

Mabtech recommends that PQ is performed at regular intervals, and the software will indicate when it is time to perform a test. When not in use, the PQ plate should be stored in its protective case in the dark, protected from dust, at room temperature.

Annual service

Mabtech offers a service package that includes an annual check-up of the reader. The performance is evaluated, and the hardware adjusted if necessary, with the consent of the user.

# 21 CFR part 11 compliance

The US (FDA) have established guidelines for the life sciences industry, known as 21 CFR part 11, respectively. These guidelines outline the use of computerized systems in clinical investigations and specify that the quality of source data obtained through these systems must match that of traditional paper records.

Mabtech Apex™ software has been designed to comply with the following guidelines in § 11.10 Controls for a closed system of 21 CFR part 11:

- A start-up test report is automatically generated each time the user starts Mabtech IRIS™ 2 and Mabtech ASTOR™ 2, and the reader goes through the start-up steps (OFF => BOOTING => READY). The user can access the report by clicking on a hyperlink in the software. The reports are also automatically saved locally at C:\ProgramData. This feature enhances the audit trail and makes it possible for users to sign off on reports where the status of the reader has been validated prior to the readout of plates.

- IQOQ and initial PQ are performed in connection with installation. Subsequent PQ can be performed using PQ Plate IRIS or PQ Plate ASTOR and the PQ function in Mabtech Apex™. A timeline is created, and baseline focus, light intensity, Spot Forming Unit (SFU) and Relative Spot Volume (RSV) values are established when registering a new PQ plate. PQ analysis is then performed regularly. Each time a PQ test is performed data is added to the timeline. The new data is compared to the baseline values, ensuring accurate reader performance over time. A PQ report outlining test criteria and results is exported in a PDF format, and the raw data from all PQ measurements are automatically saved locally at C:\ProgramData.

Clarification: Meets paragraph a) in § 11.10 Controls for closed systems.

- Copies of records are generated in both human-readable form and electronic form throughout the readout and analysis process. Data can be exported in PDF, XLSX, JPG, TIFF, PZFX, and RAW file formats.

Clarification: Meets paragraph b) in § 11.10 Controls for closed systems.

- All captured well images are saved and processed as image RAW files. These contain the untouched signal input straight from the image sensor and are free of the user bias that otherwise would be introduced by settings such as sharpness, digital gain, color, hue, and saturation. From the moment of capture, the image RAW files are never changed or altered by the software. As a result, the true RAW data is maintained and cannot be concealed by any user, including the administrator. Raw data can be automatically saved to the plate data folder after completed readout or upon cancellation of readout.

Clarification: Meets paragraph c) in § 11.10 Controls for closed systems.

- Administrator-controlled activities include:

- Admin can export an administrative audit trail, showing user activities between the start and end date (see Table 1).

- Admin can export a system audit trail showing system activities between the start and end date (see Table 1).

- Admin can export and import users.

- Admin can inactivate and reactivate users.

- Admin can generate temporary passwords and dictate user password expiration, complexity and length.

- Admin can limit the user to five failed password attempts before the user is locked out.

- Admin can set time of inactivity before the user must input their password again to continue using the system.

- Admin can restrict a user to only run selected plate templates.

- Admin can give admin privileges to created users.

- Admin can give users permission to run Performance Qualification (PQ).

- Admin defines whether the user must use Machine Default exposure values for LEDs or manual values set by admin.

- Admin determines if a user must automatically save the plate data folder after completed readout or upon cancellation of readout.

- Admin can restrict a user to only save plates using “Save as”.

- Admin can restrict a user to sign all changes before saving the plate.

Clarification: Meets paragraphs d), f), and g) in § 11.10 Controls for closed systems.

Table 1. Audit trails in Mabtech Apex™.

| Administrative audit trail | Operational audit trail | System audit trail |
| --- | --- | --- |
| 1) User created (including all settings) 2) User edited (including all settings) 3) User Log Session (date, time) 4) Password changed 5) Failed login attempts 6) User locked after too many failed password attempts and user unlocked by admin 7) Plate opened/read 8) Preset created/edited/imported/deleted 9) Admin and system audit exported 10) Readout cancelled 11) PQ performed with results 12) User edited/inactivate/reactivated/imported/exported 13) Preference settings changed 14) Templates created/edited/imported/deleted 15) Layouts imported/exported | 1) Session time (date, time, User) 2) Plate created 3) Plate saved as 4) Plate data/count data (Preset, SFU, etc.) | 1) System start-up 2) System loaded 3) System backup 4) System software crashes |

- Date, time, user, and machine identity (ID) are always generated and stored for each read and saved plate. Operational audit trails are automatically created in write-protected PDF format and specify the chronological order of all events from plate creation to last saving (see Table 1). This includes all elements that control source data: users, algorithm settings, spot counts, RSV values, time stamps, machine ID, etc. ​In addition, users can add notes to individual wells informing of mishaps in pipetting or similar. The notes are saved separately and reported within the Excel file and cannot be deleted once created. The admin can restrict a user to electronically sign all changes made to a plate and provide motivations to the distinct changes. The changes and motivations are logged in the operational audit trail.

Clarification: Meets paragraph e) in § 11.10 Controls for closed systems.

- Captured RAW images and count files generated by the RAWspot™ algorithm in Mabtech Apex™ and the audit trail of each plate are individually validated by a checksum system in the software. Upon reading, each file is automatically passed through a checksum function, and a unique block of data is generated for each file type. Upon saving, the history audit file is given a unique checksum block of data. Every time a plate is opened, the checksum validation is re-run and controlled against the stored copy, making sure that it matches the original checksum value. If any of these image RAW files, count files or history files are removed, altered, or manipulated, the checksum validation will ultimately fail when opening the plate and will provide an integrity warning. If no warnings are provided, data integrity is assured.​

Clarification: Meets paragraph h) in § 11.10 Controls for closed systems.

# Good clinical laboratory practice compatibility

Good Clinical Laboratory Practice (GCLP) is a set of standards that provides guidance on implementing Good Laboratory Practice (GLP) and Good Clinical Practice (GCP) principles to the analysis of samples from a clinical trial. GLP is a quality system that covers the organizational process and the conditions under which non-clinical laboratory studies are planned, performed, monitored, recorded, archived, and reported. It aims to promote the generation of valid, high-quality test data. The GCP principles are international, ethical, and scientific quality standards for designing, conducting, performing, monitoring, auditing, recording, analyzing, and reporting clinical trials that involve the participation of human subjects. Compliance with the practices assures that data and reported results are credible and accurate and that the rights, safety, and confidentiality of trial subjects are protected. If a customer would like to apply Mabtech IRIS™ 2 or Mabtech ASTOR™ 2 in their GCLP-regulated environment the following need to be taken into consideration:

1.	Equipment: Validation of the reader as well as installation, training, and maintenance should be completed in the specific environment for use.

2.	Standard Operating Procedures (SOPs) for handling, maintenance, etc. should be established by the user.

3.	Computer systems: Procedures for the software, such as audit trails, controlled user restrictions, and time stamps should be established.

CE mark

The readers have been tested and approved according to the EU directives for Electromagnetic compatibility and Low voltage.

# ISO 9001 and ISO 13485

The Quality Management System at Mabtech AB complies with the quality management standards ISO 9001:2015 and ISO 13485:2016. Compliance is audited annually by external audits and the certificates have been maintained without interruption since 2006 when Mabtech first aimed for certification. The scope of the certificates is as follows: Development, manufacturing, marketing, and sales of monoclonal antibodies and immunoassay instruments for in vitro applications in biomedical sciences and service of immunoassay instruments. The ISO certificates are available for download on Mabtech’s website.

Jesper Larsson
Head of Product Management 
Mabtech 2026-03-31

www.mabtech.com

| Contact Details | Mailing Address | Visiting Address | Organisation/Finance |
| --- | --- | --- | --- |
| Email: iris@mabtech.com | Mabtech AB | Mabtech AB | Organisation no: 556 276-8225 |
| Phone: +46 8 716 27 00 | Box 1233 | Augustendalstorget 9 | VAT no: SE556276822501 |
| Fax: +46 8 716 27 01 | SE-131 28 Nacka Strand | SE-131 52 Nacka Strand |  |
|  | Sweden | Sweden |  |